Calculating Optical Density Using MCiD

Andrea Vogel

Dec. 14, 2016

Overall Notes:

- Measure all regions twice (on separate days)

  - Use Pearson coefficient to determine how similar they are

  - If Pearson coefficient is 0.9, then average them together

  - If not, measure regions again and re-calculate the Pearson’s coefficient to each old sample

  - Average the samples that have Pearson’s coefficient of 0.9

- Measure all samples for given region on the same day

  - Light will change readings, will never be the same day-to-day

- Close the door to the scope room to help control for lighting differences

- Lightbox is fluorescent, so as it warms up it will change lumens

  - let it warm up fully before starting anything

- If a region is done during the morning, and you are going to complete another region during the afternoon, leave everything on while you eat luch

  - Cover the lightbox with the blue lab bench sheet

Procedure:

- If this is the first time, create a folder labeled “Your Name” on the desktop

- Open MCiD using the icon on the desktop

  - Stretch the program to cover both screens: the picture should be 1 whole screen

- Settings  Display Format  Mono (12 bit)  ok

- Settings  Input Selection  QICAM – 12 Lightbox Camera  ok

  - Check that lightbox camera is on (blue and green lights)

    - switch is on top of camera

  - Check that color camera (on top of other scope) is also on, because they are daisy-chained together

    - switch is on top of camera

- Adjust magnification and focus

  - film is on lightbox for this

  - Digitize button – looks like a video camera

    - Focus using the camera

      - use ink on film to determine if it is in focus

      - If need to adjust magnification, use the crank that the camera arm is attached to

    - has to be on digitize (dialog box open) in order to see it move

    - click ok to take image

- Establish flat field correction

  - remove film for this step

  - Change pseudocolor (grayscale bar on top menu) to SPECT2VIS

  - Adjust light intensity using intensity turner on the light box until the blank field of view looks bluish-pink

    - lumens ~1000

  - Ctrl+click on the Digitize icon  Flat Field Correction

  - If there is already a flat field correction established, click “Clear”

  - Click “Acquire” to establish a new flat field correction

    - Do not change settings!!! (but just in case, they should match the below):

      - Correction source: Digitize

      - Correction method: Pixel by pixel

      - Frame average: 4

      - Smoothing and Median filters should be clicked

  - Click “ok”

- Create a Density Calibration

  - film is back on the lightbox for this part, looking at the standards slide

  - Study type (on left-side menu)  Standards

  - Operations (on left-side menu)  Calibration

    - ROD – relative optical density

  - Calibrations (on top menu bar)  Density

    - Change to Cal Stds

      - input lowest to highest values (zero first)

    - Select a density standards file  ok

    - If you need to create a new calibration table, click “None”

    - Enter numbers from data sheet into “Standards unit” column

      - Data sheet nCi/mg x Decay rate = # entered

    - Click cursor in “Value Read” column next to first standard value

    - Digitize the standards

    - Get a zero reading (from plain film), then procedure through the standards from lightest to darkest

  - Select curve type (on dialog box)

    - cubic spline (smoothed) is usually used

      - pick the one that fits the data best (by eye)

  - Extrapolate (under curve graph)

    - will tell computer how to interpret data that falls outside the range of the standards

    - “current curve type” for both min and max

- Sampling

  - Operations  Sample

  - Have laptop nearby with Excel spreadsheet open (don’t use MCiD spreadsheet)

  - example:

| Animal | Sect 1 VP | Bkgd | Sect 2 VP | Bkgd | Sect 3 VP | Bkgd |
| --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |

  - change pseudocolor to grayscale

  - Choose whether unilateral or bilateral measuring

    - Usually use unilateral

  - Choose whether free-hand or use shape (rectangle or oval)

    - For VP, use oval

    - useful:

      - ctrl + drag = changes size

      - alt + drag = rotates

      - ctrl + click shape on menu = get measurements (to keep consistent across days/accidental changes)

      - F4 clears image, not data

  - Click on region of interest and background (2 clicks) to get data

    - if you want, hit F4 to clear the image, then re-click to make sure you are getting close numbers

    - Numbers in red are extrapolated from the standards

  - Input data into excel spreadsheet

  - Move to next slice and repeat process

  - When completely done with that animal, clear all the data using “Clear” on sheet

- Shut down

  - turn off camera on the microscope and cover the scope

  - turn off camera on the lightbox

  - put lens cover on the camera

  - wipe down the light box with sparkle and kimwipes

  - place blue lab bench sheet on the light box, soft side down

  - Shut down computer

- At end of experiment

  - get images (get help from Heather or Jinyan)

  - use micrometer to determine sizes