Comparison of DNA concentration 2X6mm vs 3X3mm Uni Core Punch

- The aim of this analysis was to compare gDNA concentration difference between 2X6mm and 3X3mm Uni Core full punch. For this comparison we used the field DBS samples collected in Zambia.

- Total number of DBS samples used for Uni Core test = 15 (8 samples = 2X6mm DBS, 7 samples = 3x3mm) and 1 negative control (2X6mm clear Whatman filter punch).

- All DBS samples used were known to be Plasmodium falciparum negative by PET-PCR. We followed our optimized protocol for gDNA extraction from DBS sample using QIAcube HT robot.

### DBS Sample gDNA Concentration (ng/µl) for 2X6mm Uni Core Punch

- gDNA concentration was assessed using HS Qubit 1x dsDNA per µL.

- The total gDNA concentration is calculated by multiplying ng/µL gDNA concentration by total elution volume (120 µL).

- We obtained median gDNA concentration = 3.55 ng/µL and range = 2.22-4.66 ng/µL. See the figure below.

### DBS Sample gDNA Concentration (ng/µl) for 3x3mm Uni Core Punch

- gDNA concentration was assessed using HS Qubit 1x dsDNA per µL.

- The total gDNA concentration is calculated by multiplying ng/µL gDNA concentration by total elution volume (120 µL).

- We obtained median gDNA concentration = 0.72 ng/µL and range = 0.44-2.01ng/µL.

- We observed a median increase of approximately 4-fold in the gDNA yield when using 2x6mm Uni Core punches compared to 3x3mm.

- There is statistically significant difference in DNA concentration obtained from 2x6mm vs 3x3mm DBS punch (P-value = 0.0003). See the figure below.

- Variation in the DNA yields between DBS punches is expected given variation in absorbance and amount of blood from field DBS puches.

- Given the significant higher DNA yield in the 2x6mm compared to the 3x3mm DBS punch using this protocol,  the use of 2x6mm is recommended.

Figure 1. Comparison of DNA concentration from 2x6mm vs 3x3mm DBS Punch.