Appendix 2: Quality control for antibiotic susceptibility testing (TLM_PRO_099)

Quality control (QC) strains to be used

| QC strain | Antibiotics to be tested | Notes |
| --- | --- | --- |
| E. coli ATCC 25922 | Disks for disk diffusion (Gram-negative) | β-lactamase negative strain Amoxicillin susceptible |
| E. coli ATCC 35218 | Β-lactam/ β-lactamase inhibitor combinations | β -lactamase positive (plasmid TEM-1, non ESBL) |
| K. pneumoniae ATCC 700603 | ESBL test | Produces β -lactamase (plasmid SHV-18, ESBL) |
| P. aeruginosa ATCC 27853 | Disks for disk diffusion (Gram-negative), E-test meropenem | Ceftazidime susceptible Meropenem susceptible |
| S. aureus ATCC 29213 | E-test vancomycin, tetracyclins, macrolides, ciprofloxacin; disks for disk diffusion (Gram-positive): only when using EUCAST | Oxacillin susceptible |
| S. aureus ATCC 25923 | Disks for disk diffusion (Gram-positive): only when using CLSI | Oxacillin susceptible |
| S. aureus ATCC 43300 | Cefoxitin disk testing | Oxacillin resistant (MRSA) |
| H. influenzae ATCC 49247 | Disks for disk diffusion | Ampicillin resistant, β -lactamase negative strain |
| S. pneumoniae ATCC 49619 | Disk for disk diffusion (Gram-positive), E-test ceftriaxone, penicillin |  |
| S. aureus ATCC BAA-976 | D-test | D-test negative |
| S. aureus ATCC BAA-977 | D-test | D-test positive |
| Salmonella Typhimurium 10433/3 | E-test azithromycin | Azithromycin resistant |

Expected results for selected QC strains

Neo-sensitab disks diameters

| Antibiotic | E. coli ATCC 25922 (mm) | P. aeruginosa ATCC 27853 (mm) | S. aureus ATCC 25923 (mm) | S. aureus ATCC 29213 (mm) | H. influenzae ATCC 49247 (mm) | S. pneumoniae ATCC 49619 (mm) |
| --- | --- | --- | --- | --- | --- | --- |
| Ampicillin 10 µg | 15 – 22 | - | 27 – 35 | - | 13 – 21 | 30 – 36 |
| Ampicillin 2 µg | - | - | - | 15 – 21 | 19 – 25 | 25 – 31 |
| Amoxi-clav 20/10 µg | 18 – 24 | - | 28 – 36 | - | 15 – 23 | - |
| Oxacillin 1 µg | - | - | 18 – 24 | - | - | < 12 |
| Penicillin 10U | - | - | 26 – 37 | - | - | 24 – 30 |
| Benzyl-penicillin 1U | - | - | - | 12 – 18 | 15 – 21 | 16 – 22 |
| Piperacillin-tazobactam 100/10 µg | 24 – 30 | 25 – 33 | - | - | - | - |
| Ciprofloxacin 5 µg | 30 – 40 | 25 – 33 | 22 – 30 | 21 – 27 | 34 – 42 | - |
| Levofloxacin 5 µg | 29 – 37 | 19 – 26 | 25 – 30 | 23 – 29 | 32 – 40 | 20 – 25 |
| Pefloxacin 5 µg | 25 – 33 | - | - | - | - | - |
| Nalidixic acid 30 µg | 22 – 28 | - | - | - | - | - |
| Norfloxacin 10 µg | - | 22 – 29 | 17 – 28 | 18 – 24 | - | - |
| Cefoxitin 30 µg | - | - | 23 – 29 | 24 – 30 | - | - |
| Cefazolin 30 µg | 21 – 27 | - | 29 – 35 | - | - | - |
| Cefuroxime 30 µg | 20 – 26 | - | - | - | - | - |
| Ceftriaxone 30 µg | 29 – 35 | 17 – 23 | 22 – 28 | - | 31 – 39 | 30 – 35 |
| Cefotaxime 30 µg | 29 – 35 | 18 – 22 | - | - | - | - |
| Ceftazidime 30 µg | 25 – 32 | 22 – 29 | - | - | - | - |
| Ceftazidime 10 µg | 23 – 29 | 21 – 27 | - | - | - | - |
| Erythromycin 15 µg | - | - | 22 – 30 | 23 – 29 | - | 25 – 30 |
| Clarithromycin 15 µg | - | - | 26 – 31 | - | 11 – 17 | 25 – 31 |
| Clindamycin 2 µg | - | - | 24 – 30 | 23 – 29 | - | 19 – 25 |
| Co-trimoxazole 1.25/23.75 µg | 23 – 29 | - | 24 – 32 | 26 – 32 | 24 – 32 | 20 – 28 |
| Nitrofurantoin 300 µg | 20 – 25 | - | 18 – 22 | - | - | - |
| Nitrofurantoin 100 µg | 17 – 23 | - | - | 17 – 23 | - | 25 – 31 |
| Fosfomycin 200 µg | 22 – 30 | - | - | - | - | - |
| Doxycycline 30 µg | 18 – 24 | - | 23 – 29 | - | - | 25 – 34 |
| Minocycline 30 µg | 19 – 25 | - | 25 – 30 | 23 – 29 | - | - |
| Tetracycline 30 µg | 18 – 25 | - | 24 – 30 | 23 – 31 | 14 – 22 | 23 – 29 |
| Chloramphenicol 30 µg | 21 – 27 | - | - | 20 – 28 | 31 – 40 | 23 – 27 |
| Gentamicin 10 µg | 19 – 26 | 17 – 23 | 19 – 27 | 19 – 25 | - | - |
| Amikacin 30 µg | 19 – 26 | 18 – 26 | 20 – 26 | 18 – 24 | - | - |
| Tobramycin 10 µg | 18 – 26 | 20 – 26 | 19 – 29 | 20 – 26 | - | - |
| Meropenem 10 µg | 28 – 34 | 27 – 33 | - | - | - | - |
| Vancomycin 30 µg | - | - | - | - | - | 20 – 27 |
| Linezolid 30 µg | - | - | 25 – 32 | - | - | 25 – 34 |
| Linezolid 10 µg | - | - | - | 21 – 27 | - | 23 – 29 |

MIC values

| Antibiotic | P. aeruginosa ATCC 27853 | S. aureus ATCC 29213 | S. pneumoniae ATCC 49619 | Salmonella Typhimurium 10433/3 |
| --- | --- | --- | --- | --- |
| Ciprofloxacin | 0.25 – 1 | 0.12 – 0.5 | - | - |
| Azithromycin | - | 0.5 – 2 | - | 16 – 64 * |
| Erythromycin | - | 0.25 – 1 | 0.03 – 0.12 | - |
| Clarithromycin | - | 0.12 – 0.5 | 0.03 – 0.12 | - |
| Doxycycline | - | 0.12 – 0.5 | 0.015 – 0.12 | - |
| Tetracycline | 8 – 32 | 0.12 – 1 | 0.06 – 0.5 | - |
| Meropenem | 0.25 – 1 | 0.03 – 0.12 | 0.03 – 0.25 | - |
| Amoxicillin/clavulanic | - | 0.12 – 0.5 | 0.03 – 0.12 | - |
| Chloramphenicol | - | 2 – 16 | 2 – 8 | - |
| Vancomycin | - | 0.5 – 2 | 0.12 – 0.5 | - |
| Penicillin | - | 0.25 – 2 | 0.25 – 1 | - |
| Ceftriaxone | 8 – 64 | 1 – 8 | 0.03 – 0.12 | - |

* strain under investigation

Beta-lactam/beta-lactam inhibitor combinations

| Antibiotic | E. coli ATCC 35218 (mm) | E. coli ATCC 35218 (MIC) |
| --- | --- | --- |
| Ampicillin 10 µg | 6 | - |
| Amoxi-clav 20/10 µg | 17 – 22 | 4 – 16 |
| Piperacillin-tazobactam 100/10 µg | 24 – 30 | 0.5 – 2 |

Detection of resistance mechanisms QC

ESBL

K. pneumoniae ATCC 700603: ESBL positive

D-test

S. aureus ATCC BAA-976: D-test negative

S. aureus ATCC BAA-977: D-test positive

MRSA

S. aureus ATCC 43300:

cefoxitin 30 µg disk diameter ≤ 21 mm

Frequency of quality control testing

- For the Unit of Tropical Laboratory Medicine: with every batch of newly arrived isolates for antibiotic susceptibility confirmation testing, at least the following QC strains are tested along:

  - E. coli ATCC 25922 (together with Gram-negative isolates),

  - S. aureus ATCC 29213 (E-tests and together with Gram-positive isolates)

  - P. aeruginosa ATCC 27853 (together with Gram-negative isolates)

  - H. influenzae ATCC 49247 (only in case Haemophilus species are tested)

- When using a new lot of antibiotic disks (Rosco Neo-Sensitabs), QC with the relevant strains should be performed:

  - E. coli ATCC 25922 for Gram-negative panel

  - E. coli ATCC 35218 (only ampicillin, amoxicillin-clavulanic acid, piperacillin-tazobactam)

  - S. aureus ATCC 29213  for Gram-positive panel (if using EUCAST)

  - S. aureus ATCC 25923 for Gram-positive panel (if using CLSI)

  - S. aureus ATCC 43300 (only cefoxitin)

  - P. aeruginosa ATCC 27853 for Gram-negative panel

  - S. pneumoniae ATCC 49619 for Gram-positive panel

  - K. pneumoniae ATCC 700603 for ESBL testing disks

  - S. aureus ATCC BAA-976 & BAA-977 for D-test

- When using a new lot of E-tests, QC with the relevant strains is performed:

  - P. aeruginosa ATCC 27853,

  - S. aureus ATCC 29213,

  - S. pneumoniae ATCC 49619

  - Salmonella Typhimurium 10433/3

- All QC strains are tested at least monthly

- In a routine laboratory, all QC strains must be tested at least on a weekly basis

QC organism maintenance:

- Culture the primary strain according to the instructions by the manufacturer. These cultures, called secondary strains or stock cultures, can be used for long term storage when stored in an appropriate way (freeze-dry, freeze between -65° and -85° C on Microbeads, cryopreservation etc). Especially for E. coli ATCC 35218 and K. pneumonia ATCC 700603, preservation at temperatures below -60°C is necessary

- When culturing these secondary strains, bench strains are obtained. Bench strains are the working cultures used for QC in the lab. They must be discarded after a week.

- Secondary strains can be used as long as they grow on subculture and as long as the tests in which they are used give good results.

- The maximum number of times a bench strain can be cultured again is three consecutive passages (see figure 1). ATCC recommends a maximum number of 5 passages from the primary strain; these 5 passages include the passage to secondary and the passage to bench strain.

3

2

1

ATCC strain: 0 passages

5 passages from ATCC strain: no more subculturing allowed

Bench strains: 2 passages

Secondary strains: 1 passage

Figure 1: subculturing from ATCC strains

QC troubleshooting (cfr CLSI M100-S28, 2018)

General concerns:

- Clavulanate is labile. Disk can lose potency, in which case an alternative lot of disks should be used.

- E. coli ATCC 35218 and K. pneumonia ATCC 700603 can appear sensitive to beta-lactam antibiotics due to spontaneous loss of the plasmid encoding the beta-lactamase.

- pH of media can be too low or too high, resulting in aberrant susceptibility to aminoglycosides, penicillins, clindamycin, macrolides, quinolones, tetracyclines. Acceptable pH range is 7.2 – 7.4. Avoid CO2 incubation, which lowers pH.

- Calcium and magnesium content of media can be too high/too low, altering susceptibility to aminoglycosides and tetracyclines. Use alternative lot of media.

- Beware of contamination when zones are too small or if there are colonies inside the zone. Subculture to determine purity and repeat if necessary.

- Inoculum too light or too heavy leads to respectively larger and smaller zones than expected. Repeat using McFarland 0.5 turbidity standard.

- Media depth can be too thin or too thick. Use agar with depth approximately 4 mm.

- Be sure to detect lighter growth in zone measurement, in case of double zone, fuzzy zone edge, …

- When bacteriostatic antibiotics are tested (such as co-trimoxazole), which can produce a “haze” of light growth, read the outer zone and disregard the haze.

References:

- CLKB SOP PRO-0289v1.1: Management and use of reference strains (09/08/2017)

- EUCAST QC tables v8.0: http://www.eucast.org/fileadmin/src/media/PDFs/EUCAST_files/QC/v_8.0_EUCAST_QC_tables_routine_and_extended_QC.pdf

- Performance Standards for antimicrobial susceptibility testing; Twenty-Eighth Informational Supplement. CLSI document M100-S28 . CLSI, 950 West Valley Road, Suite 2500, Wayne, Pennsylvania 19087 USA, 2018. Available from http://em100.edaptivedocs.info/Login.aspx

- American Type Culture Collection. Reference strains: how many passages are too many? Technical bulletin no. 6; 2013. Available from https://www.atcc.org/~/media/PDFs/Technical%20Bulletins/tb06.ashx

| Revision | Changes with respect to the previous version |
| --- | --- |
| 15/02/2018 | Initial version |
| 13/07/2018 | Maintenance of QC strains: Tertiary strains represented as blood agars, not tubes |
| 10/01/2019 | Tertiary strains replaced by bench strains |

| Name and function | Date | Signature |
| --- | --- | --- |
| Author |  |  |
| Sien Ombelet | 30/01/2018 |  |
| Reviewed by |  |  |
| Jan Jacobs | 06/02/2018 | See email 180206 SOP AST |
| Approved by |  |  |
| Jan Jacobs | 13/02/2018 | Unit Meeting 180212 |