## General
Hg | Primary Mercury Gas Standard | (PMGS)
Date: | 2022-07-07
Operator: | RZ
Analyser: | PSA SG II
Analysis method: | F5min measureAR10
PMGS/Reference Standard: | VSL
Candidate generator: | VSL
Range: | 0.9 – 2.5 ug/m3
Project: | SI-Hg
Remarks:
Instruction: | PMGS sheet for data from primary mercury gas standard
Bell-jar sheet for data from bell-jar injections
Unknown sheet for data from unknown mercury gas generator
Summary sheet to obtain a table for CurveFit calculation
Flow sheet for data from the calibration of the MFC with the BIOS
Fill out the green cells
Blue cells can be used as input for CurveFit
Version: | Description | Date
1 | New template | 2023-02-09

## Candidate G1
Generator: | 1 | 1,2 or 3 |  |  | Remarks:
Diffusion cells: | 33 | numbers
Mercury diffusion: | 10.52 | ng min-1
u mercury diffusion: | 0.3156 | ng min-1 | k = 1
Analyser: | PSA SG II
Measurement uncertainty: | 1.5 | % | k = 1
Sample time: | 5 | Min
Sample flow: | 0.6 | L min-1
Total flow: | 3 | L
Sampled via PSA stream selector
Setpoint: |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Set point MFC nitrogen: | L min-1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Set point MFC air: | L min-1 | 3 | 3.5 | 4.15 | 5 | 6.3 | 8.25
Total system flow: | L min-1 | 4.1047 | 4.649 | 5.3566 | 6.2819 | 7.6971 | 9.8199
Mercury concentration: | ng m-3 | 2562.9122 | 2262.8455 | 1963.9271 | 1674.6424 | 1366.7415 | 1071.2895
U mercury concentration (k = 2): | ng m-3
U mercury concentration (k = 2): | %

## Flow G1
MFC: |  | Air diffusion chamber 1 |  | MFC: |  | Nitrogen diffusion chamber 1
Range: |  | 0.5-25 | L min-1 | Range: |  | 0.5 | L min-1
Calibration date: |  | 2022-03-28 |  | Calibration date: |  | 2022-03-28
Data file: |  | K:\Beheer\Hg\MFC\2022-03-28\2022-03-28 Template_Calibration Hg MFCs.xlsx |  | Data file: |  | K:\Beheer\Hg\MFC\2022-03-28\2022-03-28 Template_Calibration Hg MFCs.xlsx
Certificate: |  | - |  | Certificate: |  | -
Calibration coeficients |  |  |  | Calibration coeficients
Model equation y = a[0] + a[1]*x |  |  |  | Set point MFC: | 500 | ml min-1
Regression coefficients: |  |  |  | Read out MFM: | 525.22 | ml min-1
a[0] | 0.31364 | 0.0027834
a[1] | 1.0886 | 0.00072345

## PMGS G2
Generator: | 2 | 1,2 or 3 |  |  | Remarks:
Diffusion cells: | 18, 19, 20 | numbers
Mercury diffusion: | 4.03 | ng min-1
u mercury diffusion: | 0.12 | ng min-1 | k = 1
Analyser: | PSA SG II
Measurement uncertainty: | 1.5 | % | k = 1
Sample time: | 5 | Min
Sample flow: | 0.6 | L min-1
Total flow: | 3 | L
Sampled via PSA stream selector
Setpoint: |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Set point MFC nitrogen: | L min-1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Set point MFC air: | L min-1 | 0.74 | 1.01 | 1.37 | 1.93 | 2.8 | 0.74
Total system flow: | L min-1 | 1.8102 | 2.1136 | 2.5175 | 3.1442 | 4.1144 | 1.8102
Mercury concentration: | ng m-3 | 2226.3174 | 1906.7006 | 1600.8039 | 1281.7121 | 979.4751 | 2226.3174
U mercury concentration (k = 2): | ng m-3 | 111.3159 | 95.335 | 80.0402 | 64.0856 | 48.9738 | 111.3159
U mercury concentration (k = 2): | % | bar: 1 to 6

## Flow G2
MFC: |  | Air diffusion chamber 2 |  | MFC: |  | Nitrogen diffusion chamber 2
Range: |  | 0.5 - 25 | L min-1 | Range: |  | 0.5 | L min-1
Calibration date: |  | 2022-03-28 |  | Calibration date: |  | 2022-03-28
Data file: |  | K:\Beheer\Hg\MFC\2022-03-28\2022-03-28 Template_Calibration Hg MFCs.xlsx |  | Data file: |  | K:\Beheer\Hg\MFC\2022-03-28\2022-03-28 Template_Calibration Hg MFCs.xlsx
Certificate: |  | - |  | Certificate: |  | -
Calibration coeficients |  |  |  | Calibration coeficients
Model equation y = a[0] + a[1]*x + a[2]*x^2 + a[3]*x^3 |  |  |  | Set point MFC: | 846 | ml min-1
Regression coefficients: |  |  |  | Read out MFM: | 501.4 | ml min-1
a[0] | 0.47476 | 0.0073609
a[1] | 1.1295 | 0.0044697
a[2] | -0.0033275 | 0.0005749
a[3] | 8.7063e-05 | 1.8028e-05

## Data
| DateTime | Pk Area | Name | Setpoint_candidate | Valid
0 | 2022-07-07 17:16 | 131.922 | Sample Stream 3A Zero Air | 0 | TRUE
1 | 2022-07-07 17:21 | 138.713 | Sample Stream 3B Zero Air | 0 | TRUE
2 | 2022-07-07 17:26 | 118.306 | Sample Stream 3A Zero Air | 0 | TRUE
3 | 2022-07-07 17:32 | 106.723 | Sample Stream 3B Zero Air | 0 | TRUE
4 | 2022-07-07 17:37 | 86.16 | Sample Stream 3A Zero Air | 0 | TRUE
5 | 2022-07-07 17:42 | 82.738 | Sample Stream 3B Zero Air | 0 | TRUE
6 | 2022-07-07 17:48 | 67.175 | Sample Stream 3A Zero Air | 0 | TRUE
7 | 2022-07-07 17:53 | 66.993 | Sample Stream 3B Zero Air | 0 | TRUE
8 | 2022-07-07 17:58 | 50.557 | Sample Stream 3A Zero Air | 0 | TRUE
9 | 2022-07-07 18:03 | 56.968 | Sample Stream 3B Zero Air | 0 | TRUE
10 | 2022-07-07 18:09 | 43.531 | Sample Stream 3A Zero Air | 0 | TRUE
11 | 2022-07-07 18:20 | 10.725 | Sample Stream 3B Zero Air | 0 | TRUE
12 | 2022-07-07 18:25 | 5707.606 | Sample Stream 1A 3x3mm | 2563 | TRUE
13 | 2022-07-07 18:31 | 5705.001 | Sample Stream 1B 3x3mm | 2563 | TRUE
14 | 2022-07-07 18:36 | 6158.898 | Sample Stream 2A 1x 8mm | 2563 | TRUE
15 | 2022-07-07 18:41 | 6103.03 | Sample Stream 2B 1x 8mm | 2563 | TRUE
16 | 2022-07-07 18:47 | 5576.463 | Sample Stream 1A 3x3mm | 2563 | TRUE
17 | 2022-07-07 18:52 | 5558.723 | Sample Stream 1B 3x3mm | 2563 | TRUE
18 | 2022-07-07 18:57 | 5921.83 | Sample Stream 2A 1x 8mm | 2563 | TRUE
19 | 2022-07-07 19:03 | 6026.857 | Sample Stream 2B 1x 8mm | 2563 | TRUE
20 | 2022-07-07 19:08 | 5686.081 | Sample Stream 1A 3x3mm | 2563 | TRUE
21 | 2022-07-07 19:13 | 5497.614 | Sample Stream 1B 3x3mm | 2563 | TRUE
22 | 2022-07-07 19:18 | 5966.472 | Sample Stream 2A 1x 8mm | 2563 | TRUE
23 | 2022-07-07 19:24 | 6104.256 | Sample Stream 2B 1x 8mm | 2563 | TRUE
24 | 2022-07-07 19:29 | 5636.063 | Sample Stream 1A 3x3mm | 2563 | TRUE
25 | 2022-07-07 19:34 | 5681.568 | Sample Stream 1B 3x3mm | 2563 | TRUE
26 | 2022-07-07 19:40 | 6065.253 | Sample Stream 2A 1x 8mm | 2563 | TRUE
27 | 2022-07-07 19:45 | 6089.427 | Sample Stream 2B 1x 8mm | 2563 | TRUE
28 | 2022-07-07 19:50 | 5648.593 | Sample Stream 1A 3x3mm | 2563 | TRUE
29 | 2022-07-07 19:56 | 5564.529 | Sample Stream 1B 3x3mm | 2563 | TRUE
30 | 2022-07-07 20:01 | 6012.301 | Sample Stream 2A 1x 8mm | 2563 | TRUE
31 | 2022-07-07 20:06 | 6133.661 | Sample Stream 2B 1x 8mm | 2563 | TRUE
32 | 2022-07-07 20:11 | 5692.174 | Sample Stream 1A 3x3mm | 2563 | TRUE
33 | 2022-07-07 20:17 | 5672.234 | Sample Stream 1B 3x3mm | 2563 | TRUE
34 | 2022-07-07 20:22 | 5958.714 | Sample Stream 2A 1x 8mm | 2563 | TRUE
35 | 2022-07-07 20:27 | 5946.525 | Sample Stream 2B 1x 8mm | 2563 | TRUE
36 | 2022-07-08 08:33 | 4808.437 | Sample Stream 1A 3x3mm | 2263 | TRUE
37 | 2022-07-08 08:38 | 4895.792 | Sample Stream 1B 3x3mm | 2263 | TRUE
38 | 2022-07-08 08:44 | 5418.217 | Sample Stream 2A 1x 8mm | 2263 | TRUE
39 | 2022-07-08 08:49 | 5416.36 | Sample Stream 2B 1x 8mm | 2263 | TRUE
40 | 2022-07-08 08:54 | 4823.206 | Sample Stream 1A 3x3mm | 2263 | TRUE
41 | 2022-07-08 09:00 | 4934.479 | Sample Stream 1B 3x3mm | 2263 | TRUE
42 | 2022-07-08 09:05 | 5360.294 | Sample Stream 2A 1x 8mm | 2263 | TRUE
43 | 2022-07-08 09:10 | 5332.288 | Sample Stream 2B 1x 8mm | 2263 | TRUE
44 | 2022-07-08 09:16 | 4814.217 | Sample Stream 1A 3x3mm | 2263 | TRUE
45 | 2022-07-08 09:21 | 4906.19 | Sample Stream 1B 3x3mm | 2263 | TRUE
46 | 2022-07-08 09:26 | 5280.131 | Sample Stream 2A 1x 8mm | 2263 | TRUE
47 | 2022-07-08 09:31 | 5323.75 | Sample Stream 2B 1x 8mm | 2263 | TRUE
48 | 2022-07-08 09:37 | 4857.369 | Sample Stream 1A 3x3mm | 2263 | TRUE
49 | 2022-07-08 09:42 | 4871.329 | Sample Stream 1B 3x3mm | 2263 | TRUE
50 | 2022-07-08 09:47 | 5331.534 | Sample Stream 2A 1x 8mm | 2263 | TRUE
51 | 2022-07-08 09:53 | 5335.218 | Sample Stream 2B 1x 8mm | 2263 | TRUE
52 | 2022-07-08 09:58 | 4847.633 | Sample Stream 1A 3x3mm | 2263 | TRUE
53 | 2022-07-08 10:03 | 4896.84 | Sample Stream 1B 3x3mm | 2263 | TRUE
54 | 2022-07-08 10:08 | 5276.741 | Sample Stream 2A 1x 8mm | 2263 | TRUE
55 | 2022-07-08 10:14 | 5384.726 | Sample Stream 2B 1x 8mm | 2263 | TRUE
56 | 2022-07-08 10:19 | 4806.056 | Sample Stream 1A 3x3mm | 2263 | TRUE
57 | 2022-07-08 10:24 | 4836.864 | Sample Stream 1B 3x3mm | 2263 | TRUE
58 | 2022-07-08 10:30 | 5262.628 | Sample Stream 2A 1x 8mm | 2263 | TRUE
59 | 2022-07-08 10:35 | 5343.428 | Sample Stream 2B 1x 8mm | 2263 | TRUE
60 | 2022-07-08 10:40 | 4765.716 | Sample Stream 1A 3x3mm | 2263 | TRUE
61 | 2022-07-08 10:45 | 4795.068 | Sample Stream 1B 3x3mm | 2263 | TRUE
62 | 2022-07-08 10:51 | 5266.995 | Sample Stream 2A 1x 8mm | 2263 | TRUE
63 | 2022-07-08 10:56 | 5334.762 | Sample Stream 2B 1x 8mm | 2263 | TRUE
64 | 2022-07-08 11:44 | 4122.507 | Sample Stream 1A 3x3mm | 1964 | TRUE
65 | 2022-07-08 11:49 | 4111.44 | Sample Stream 1B 3x3mm | 1964 | TRUE
66 | 2022-07-08 11:54 | 4653.122 | Sample Stream 2A 1x 8mm | 1964 | TRUE
67 | 2022-07-08 12:00 | 4694.176 | Sample Stream 2B 1x 8mm | 1964 | TRUE
68 | 2022-07-08 12:05 | 4120.091 | Sample Stream 1A 3x3mm | 1964 | TRUE
69 | 2022-07-08 12:10 | 4150.341 | Sample Stream 1B 3x3mm | 1964 | TRUE
70 | 2022-07-08 12:16 | 4660.818 | Sample Stream 2A 1x 8mm | 1964 | TRUE
71 | 2022-07-08 12:21 | 4687.277 | Sample Stream 2B 1x 8mm | 1964 | TRUE
72 | 2022-07-08 12:26 | 4132.367 | Sample Stream 1A 3x3mm | 1964 | TRUE
73 | 2022-07-08 12:31 | 4077.464 | Sample Stream 1B 3x3mm | 1964 | TRUE
74 | 2022-07-08 12:37 | 4598.19 | Sample Stream 2A 1x 8mm | 1964 | TRUE
75 | 2022-07-08 12:42 | 4648.892 | Sample Stream 2B 1x 8mm | 1964 | TRUE
76 | 2022-07-08 12:47 | 4092.639 | Sample Stream 1A 3x3mm | 1964 | TRUE
77 | 2022-07-08 12:53 | 4109.307 | Sample Stream 1B 3x3mm | 1964 | TRUE
78 | 2022-07-08 12:58 | 4627.268 | Sample Stream 2A 1x 8mm | 1964 | TRUE
79 | 2022-07-08 13:03 | 4694.673 | Sample Stream 2B 1x 8mm | 1964 | TRUE
80 | 2022-07-08 13:08 | 4135.27 | Sample Stream 1A 3x3mm | 1964 | TRUE
81 | 2022-07-08 13:14 | 4087.522 | Sample Stream 1B 3x3mm | 1964 | TRUE
82 | 2022-07-08 13:19 | 4624.017 | Sample Stream 2A 1x 8mm | 1964 | TRUE
83 | 2022-07-08 13:24 | 4739.737 | Sample Stream 2B 1x 8mm | 1964 | TRUE
84 | 2022-07-08 13:30 | 4163.665 | Sample Stream 1A 3x3mm | 1964 | TRUE
85 | 2022-07-08 13:35 | 4119.781 | Sample Stream 1B 3x3mm | 1964 | TRUE
86 | 2022-07-08 13:40 | 4669.307 | Sample Stream 2A 1x 8mm | 1964 | TRUE
87 | 2022-07-08 13:46 | 4639.856 | Sample Stream 2B 1x 8mm | 1964 | TRUE
88 | 2022-07-11 10:54 | 3278.315 | Sample Stream 1A 3x3mm | 1675 | TRUE
89 | 2022-07-11 10:59 | 3243.008 | Sample Stream 1B 3x3mm | 1675 | TRUE
90 | 2022-07-11 11:05 | 3959.265 | Sample Stream 2A 1x 8mm | 1675 | TRUE
91 | 2022-07-11 11:10 | 4001.233 | Sample Stream 2B 1x 8mm | 1675 | TRUE
92 | 2022-07-11 11:15 | 3280.2 | Sample Stream 1A 3x3mm | 1675 | TRUE
93 | 2022-07-11 11:20 | 3296.37 | Sample Stream 1B 3x3mm | 1675 | TRUE
94 | 2022-07-11 11:26 | 3867.467 | Sample Stream 2A 1x 8mm | 1675 | TRUE
95 | 2022-07-11 11:31 | 3880.032 | Sample Stream 2B 1x 8mm | 1675 | TRUE
96 | 2022-07-11 11:36 | 3209.86 | Sample Stream 1A 3x3mm | 1675 | TRUE
97 | 2022-07-11 11:42 | 3219.361 | Sample Stream 1B 3x3mm | 1675 | TRUE
98 | 2022-07-11 11:47 | 3851.901 | Sample Stream 2A 1x 8mm | 1675 | TRUE
99 | 2022-07-11 11:52 | 3906.197 | Sample Stream 2B 1x 8mm | 1675 | TRUE
100 | 2022-07-11 11:58 | 3219.079 | Sample Stream 1A 3x3mm | 1675 | TRUE
101 | 2022-07-11 12:03 | 3210.567 | Sample Stream 1B 3x3mm | 1675 | TRUE
102 | 2022-07-11 12:08 | 3869.273 | Sample Stream 2A 1x 8mm | 1675 | TRUE
103 | 2022-07-11 12:13 | 3916.681 | Sample Stream 2B 1x 8mm | 1675 | TRUE
104 | 2022-07-11 12:40 | 2449.023 | Sample Stream 1A 3x3mm | 1367 | TRUE
105 | 2022-07-11 12:45 | 2455.219 | Sample Stream 1B 3x3mm | 1367 | TRUE
106 | 2022-07-11 12:51 | 3191.793 | Sample Stream 2A 1x 8mm | 1367 | TRUE
107 | 2022-07-11 12:56 | 3184.948 | Sample Stream 2B 1x 8mm | 1367 | TRUE
108 | 2022-07-11 13:01 | 2419.057 | Sample Stream 1A 3x3mm | 1367 | TRUE
109 | 2022-07-11 13:06 | 2452.825 | Sample Stream 1B 3x3mm | 1367 | TRUE
110 | 2022-07-11 13:12 | 3152.375 | Sample Stream 2A 1x 8mm | 1367 | TRUE
111 | 2022-07-11 13:17 | 3163.291 | Sample Stream 2B 1x 8mm | 1367 | TRUE
112 | 2022-07-11 13:22 | 2411.013 | Sample Stream 1A 3x3mm | 1367 | TRUE
113 | 2022-07-11 13:28 | 2454.915 | Sample Stream 1B 3x3mm | 1367 | TRUE
114 | 2022-07-11 13:33 | 3172.703 | Sample Stream 2A 1x 8mm | 1367 | TRUE
115 | 2022-07-11 13:38 | 3170.711 | Sample Stream 2B 1x 8mm | 1367 | TRUE
116 | 2022-07-11 13:43 | 2378.216 | Sample Stream 1A 3x3mm | 1367 | TRUE
117 | 2022-07-11 13:49 | 2413.553 | Sample Stream 1B 3x3mm | 1367 | TRUE
118 | 2022-07-11 13:54 | 3166.323 | Sample Stream 2A 1x 8mm | 1367 | TRUE
119 | 2022-07-11 13:59 | 3108.629 | Sample Stream 2B 1x 8mm | 1367 | TRUE
120 | 2022-07-11 14:05 | 2401.556 | Sample Stream 1A 3x3mm | 1367 | TRUE
121 | 2022-07-11 14:10 | 2397.692 | Sample Stream 1B 3x3mm | 1367 | TRUE
122 | 2022-07-11 14:15 | 3080.187 | Sample Stream 2A 1x 8mm | 1367 | TRUE
123 | 2022-07-11 14:21 | 3148.409 | Sample Stream 2B 1x 8mm | 1367 | TRUE
124 | 2022-07-11 14:26 | 2417.945 | Sample Stream 1A 3x3mm | 1367 | TRUE
125 | 2022-07-11 14:31 | 2413.266 | Sample Stream 1B 3x3mm | 1367 | TRUE
126 | 2022-07-11 14:36 | 3103.321 | Sample Stream 2A 1x 8mm | 1367 | TRUE
127 | 2022-07-11 14:42 | 3141.286 | Sample Stream 2B 1x 8mm | 1367 | TRUE
128 | 2022-07-11 15:29 | 5418.07 | Sample Stream 1A 3x3mm | 1071 | TRUE
129 | 2022-07-11 15:35 | 5408.655 | Sample Stream 1B 3x3mm | 1071 | TRUE
130 | 2022-07-11 15:40 | 2490.663 | Sample Stream 2A 1x 8mm | 1071 | TRUE
131 | 2022-07-11 15:45 | 2483.261 | Sample Stream 2B 1x 8mm | 1071 | TRUE
132 | 2022-07-11 15:51 | 5353.008 | Sample Stream 1A 3x3mm | 1071 | TRUE
133 | 2022-07-11 15:56 | 5452.835 | Sample Stream 1B 3x3mm | 1071 | TRUE
134 | 2022-07-11 16:01 | 2470.221 | Sample Stream 2A 1x 8mm | 1071 | TRUE
135 | 2022-07-11 16:06 | 2524.515 | Sample Stream 2B 1x 8mm | 1071 | TRUE
136 | 2022-07-11 16:12 | 5420.539 | Sample Stream 1A 3x3mm | 1071 | TRUE
137 | 2022-07-11 16:17 | 5386.831 | Sample Stream 1B 3x3mm | 1071 | TRUE
138 | 2022-07-11 16:22 | 2432.525 | Sample Stream 2A 1x 8mm | 1071 | TRUE
139 | 2022-07-11 16:28 | 2558.127 | Sample Stream 2B 1x 8mm | 1071 | TRUE
140 | 2022-07-11 16:33 | 5425.961 | Sample Stream 1A 3x3mm | 1071 | TRUE
141 | 2022-07-11 16:38 | 5409.275 | Sample Stream 1B 3x3mm | 1071 | TRUE
142 | 2022-07-11 16:44 | 2449.298 | Sample Stream 2A 1x 8mm | 1071 | TRUE
143 | 2022-07-11 16:49 | 2495.968 | Sample Stream 2B 1x 8mm | 1071 | TRUE
144 | 2022-07-11 16:54 | 5402.566 | Sample Stream 1A 3x3mm | 1071 | TRUE
145 | 2022-07-11 16:59 | 5058.826 | Sample Stream 1B 3x3mm | 1071 | TRUE
146 | 2022-07-11 17:05 | 2511.581 | Sample Stream 2A 1x 8mm | 1071 | TRUE
147 | 2022-07-11 17:10 | 2476.25 | Sample Stream 2B 1x 8mm | 1071 | TRUE
148 | 2022-07-11 17:15 | 5418.577 | Sample Stream 1A 3x3mm | 1071 | TRUE
149 | 2022-07-11 17:21 | 5508.976 | Sample Stream 1B 3x3mm | 1071 | TRUE
150 | 2022-07-11 17:26 | 2447.468 | Sample Stream 2A 1x 8mm | 1071 | TRUE
151 | 2022-07-11 17:31 | 2503.87 | Sample Stream 2B 1x 8mm | 1071 | TRUE
152 | 2022-07-11 18:34 | 53.109 | Sample Stream 3A Zero Air | 0 | TRUE
153 | 2022-07-11 18:39 | 35.405 | Sample Stream 3B Zero Air | 0 | TRUE
154 | 2022-07-11 18:44 | 37.919 | Sample Stream 3A Zero Air | 0 | TRUE
155 | 2022-07-11 18:50 | 27.039 | Sample Stream 3B Zero Air | 0 | TRUE
156 | 2022-07-11 18:55 | 26.797 | Sample Stream 3A Zero Air | 0 | TRUE
157 | 2022-07-11 19:00 | 17.843 | Sample Stream 3B Zero Air | 0 | TRUE
158 | 2022-07-11 19:05 | 20.219 | Sample Stream 3A Zero Air | 0 | TRUE
159 | 2022-07-11 19:11 | 14.23 | Sample Stream 3B Zero Air | 0 | TRUE
160 | 2022-07-11 19:16 | 12.152 | Sample Stream 3A Zero Air | 0 | TRUE
161 | 2022-07-11 19:21 | 14.661 | Sample Stream 3B Zero Air | 0 | TRUE
162 | 2022-07-11 19:27 | 8.857 | Sample Stream 3A Zero Air | 0 | TRUE
163 | 2022-07-11 19:32 | 9.164 | Sample Stream 3B Zero Air | 0 | TRUE
164 | 2022-07-11 19:37 | 10.4 | Sample Stream 3A Zero Air | 0 | TRUE
165 | 2022-07-11 19:42 | 6.653 | Sample Stream 3B Zero Air | 0 | TRUE
166 | 2022-07-11 19:48 | 5.904 | Sample Stream 3A Zero Air | 0 | TRUE
167 | 2022-07-11 19:53 | 27.959 | Sample Stream 3B Zero Air | 0 | TRUE
168 | 2022-07-11 19:58 | 6.065 | Sample Stream 3A Zero Air | 0 | TRUE
169 | 2022-07-11 20:04 | 5.917 | Sample Stream 3B Zero Air | 0 | TRUE
170 | 2022-07-11 20:09 | 3.488 | Sample Stream 3A Zero Air | 0 | TRUE
171 | 2022-07-11 20:14 | 7.167 | Sample Stream 3B Zero Air | 0 | TRUE
172 | 2022-07-11 20:20 | 4.254 | Sample Stream 3A Zero Air | 0 | TRUE
173 | 2022-07-11 20:25 | 6.371 | Sample Stream 3B Zero Air | 0 | TRUE
174 | 2022-07-11 20:30 | 8.253 | Sample Stream 3A Zero Air | 0 | TRUE
175 | 2022-07-11 20:35 | 5.577 | Sample Stream 3B Zero Air | 0 | TRUE
176 | 2022-07-11 20:41 | 5.685 | Sample Stream 3A Zero Air | 0 | TRUE
177 | 2022-07-11 20:46 | 4.755 | Sample Stream 3B Zero Air | 0 | TRUE
178 | 2022-07-11 20:51 | 5.069 | Sample Stream 3A Zero Air | 0 | TRUE
179 | 2022-07-11 20:57 | 6.152 | Sample Stream 3B Zero Air | 0 | TRUE
180 | 2022-07-11 21:02 | 5.905 | Sample Stream 3A Zero Air | 0 | TRUE
181 | 2022-07-11 21:07 | 5.693 | Sample Stream 3B Zero Air | 0 | TRUE