Evaluation of three different mechanical chest compression devices
– protocol for a prospective, simulation based, mannequin study

Mads Kjeldsen BSc,1 Rune Pedersen BSc,1 Simon Ladehoff Thomsen MD,2,3 Fritz Møller PM,3 Peter Brian Madsen Burholt PM,3 Jacob Thornvig Andersen PM,3 Niels Ole Saaby Hald MD,1,3 Bodil Steen Rasmussen MD PhD,1,2 Jannie Kristine Bang Gram PhD,3 and Thomas Lass Klitgaard MD PhD2*

aa

1 Department of Clinical Medicine, Aalborg University, Aalborg, Denmark

2 Department of Anaesthesia and Intensive Care, Aalborg University Hospital, Aalborg, Denmark

3 Emergency Medical Services, North Denmark Region, Aalborg, Denmark

*Corresponding author. Email: tlk@rn.dk; mobile: +45 20921197; Aalborg University Hospital, Hobrovej 18-22, 9000 Aalborg, Denmark.

Introduction

Around 5000 out-of-hospital cardiac arrests (OHCA) occur each year in Denmark, and the overall chance of survival is low.1 High quality chest compression is a corner stone in the resuscitative effort of cardiac arrest for obtaining sustained return of spontaneous circulation (ROSC) (e.g. spontaneous movement, coughing, breathing, a palpable pulse or a measurable blood pressure)  for at least 20 minutes) and maximizing chances of survival,2 and consists of maintaining optimal chest compression rate and depth, and minimising “hands-off time” to ensure the highest possible quality of circulatory support during low-flow time (i.e., time of resuscitation with chest compressions).3,4 Observational data demonstrate progressive deterioration over time of the cardio-pulmonary resuscitation (CPR) effort when performed manually.5 Further, CPR-quality may be compromised due to several reasons such as lack of manual resources, fatigue, and competing task such as defibrillation, airway management, and administration of medicine. With application of a mechanical chest compression device, vascular flow parameters, e.g. mean arterial blood pressure, may improve.6 Additionally, manual CPR in a moving vehicle or during patient transport may be impeded compared to employing a mechanical device,7–14 and represents a significant security risk for the ambulance personnel.

Currently, a number of different mechanical chest compression systems have been devised and are in clinical use: e.g. LUCAS III (Stryker Medical), Autopulse (Zoll Medical), and the recently developed Corpuls system (GS). These machines employ different compression mechanisms; e.g. compression bands or pistons. Such devices never tire, deliver consistent compressions, but are bulky, and can present difficulty in practical application, e.g. due to patient size and the effort required with attaching the to the patient.15 In simulation settings, mechanical devices have been demonstrated to provide more efficient chest compression compared to manual technique,16,17 but with poorer outcome in animal studies.18 In both a mechanical and an animal model, the Corpuls device has been demonstrated to provide superior flow compared to the LUCAS II.19,20 However, the use of mechanical compression is associated with increased risk of internal organ injury, increasing with the longer duration of chest compressions.21 One randomized trial of mechanical versus manual chest compression technique found worse outcome in the group allocated to mechanical compressions,5 and another trial failed to demonstrate superiority of mechanical over manual compression technique.22 However, the overall evidence on the superiority of either approach is thus far inconclusive,4 though manual compressions appear superior to the use of the Autopulse device and equal to the LUCAS II, but with a lower risk of pneumothorax or haematoma.23 However, the Autopulse system does not deliver CPR in accordance with the current European Resuscitation Council guidelines, e.g. with a lower-than-recommended number of compressions per minute.24 In addition, the COMPRESS prospective study aims to assess the effectiveness of the Corpuls system in real-life setting, but results are pending.25 Overall, choice of mechanical chest compression device is still a matter of debate.

We will conduct this study to compare the ease of use and the quality of the delivered chest compressions in a simulation setting, comparing three different mechanical chest compression devices; the LUCAS III, Autopulse CPR, and Corpuls devices.

Methods

A total of 10-20 paramedics will test the following mechanical chest compression devices:

- Corpuls CPR (GS Elektromedizinische Geräte G. Stemple GmbH, Kaufering, Germany)

- AutoPulse Resuscitation System (ZOLL Circulation, Sunnyvale, California)

- Lucas III (Physio-Control Sweden/Jolife AB, Lund, Sweden).

We will use a mannequin setting using the AmbuMan Advanced (Ambu A/S, Denmark) for performing CPR.

All participants will be trained in the use of all three devices prior to completion of the simulation study.

Baseline characteristics af participants

To describe the participants in the study, we will register the following characteristics: 1) age; 2) sex; 3) height; 4) weight; 5) years of experience as a PM/anaesthesiologist work in the prehospital setting; 6) previous experience with either device (self-evaluated competence: 0-10 (0 not at all, 10 very).

Scenarios

We will conduct the following four scenarios.

- The mannequin is placed on the floor with one assistant performing manual chest compressions. The participant will singlehandedly apply the compression device. 6 minutes of CPR will be performed in each instance.

- Transfer from 1st floor to ground level. The mannequin will be transferred from the 1st floor to ground level using the stairs, onto the ambulance gurney, and subsequently into the ambulance. Initially two minutes of CPR on the floor, transfer via stairs to the ambulance gurney, followed by four minutes of CPR on the ambulance gurney. For each compression device we will use the specific transfer device.

- Pack-down. Removal, cleaning, and packaging of the chest compression device to emulate preparation for a new dispatch. The participant will remove the compression device from the mannequin, clean it, and pack the device to be ready for a new mission.

- Transfer via ambulance. We will use a standardized road test with both acceleration, deceleration, left and right turns.

Each participant will complete scenarios 1 through 3 once per device. We will randomise the sequence of chest compression devices used by each participant using opaque envelopes.

Scenario 4 will be completed the same number of times per device as the number of participants.

Recommendations for compressions are to follow ERC-guidelines:3

Participants will be instructed in following the recommendations for chest compressions as specified by the ERC: rate of 100-120 compressions per minute; compression ratio of 30:2 between compressions and ventilations; compression to 5-6 cm of depth; to allow full recoil of the chest wall; and with minimal interruption of compressions.

Outcomes

Quality of chest compressions:

- Fraction of correctly delivered compression depth (target of 5 to 6 cm, both included).

  - Fraction of insufficient compression depth (less than 5 cm).

  - Fraction of excessive compression depth (more than 6 cm).

- Compression rate as number of compression per minute (100 to 120 compressions per minute).

- Hands-off-time as the percentage of time chest compressions are performed out of total CPR-time.

- Fraction of compressions with correct placement.

Data for evaluation of the above mentioned outcomes will be captured directly from the AmbuMan Advanced mannequin.

Time interval measurements:

- Scenario 1:

  - From start of scenario until arrival at patient.

  - From arrival at patient until device is attached to the mannequin and operational.

  - Total elapsed time.

- Scenario 2:

  - From pickup of the mannequin until it is loaded into the ambulance.

- Scenario 3:

  - Total elapsed time.

- Scenario 4:

  - Total elapsed time.

Correct placement of device upon first try:

- Percentage of devices correctly placed at first try as evaluated by observer. Will be evaluated after scenario 1.

Device movement during transfer:

- Percentage of devices that moved during patient transfer. Will be evaluated after scenarios 1, 2 and 4.

- Number of re-placements of device during transfer. Will be evaluated after scenario 2.

- Maximum movement of device during transfer. Will be evaluated after scenarios 1 and 4.

Usability

For each device we will evaluate the usability by asking a list of predefined questions. We will also evaluate the overall satisfaction with the device via a questionary.

Statistics

We will present count variables as numbers and percentages and continuous variables as means and standard deviations. To analyse a group of either measurement times or questionnaire scores, we will use a mixed restricted maximum likelihood (REM) regression with a random intercept at the lowest hierarchical level, person ID. The dataset is in the long format with either scenarios or questions as a categorical variable. We will estimate separate random intercept variances by each scenario or question. We will combine fractions and the compression rate in a mixed Poisson regression with a random intercept by the person ID. The dataset is in a long format with a categorical variable for each type of measurement. We will report means with 95% confidence intervals by each compression device. A two-sided p-value of less than 0.05 will be considered statistically significant in any analysis. All analyses are blinded to the device allocation and will be conducted using either R or STATA.

Timeline

November-February 2023: Drafting of protocol

February-May 2024: Funding

May-September 2024: Data collection

September-November 2024: Data analysis and writing of scientific report

Discussion

We will use simulations based on real-life scenarios to emulate such conditions. However, this study will only consider the quality of provided CPR and the ease of use of the different devices, but not on conjunction to other interventions relevant to the acute care of patients with cardiac arrest, e.g. airway management, emergency thoracotomies, medication etc. The simulation setup will allow a direct comparison between the different devices with minimal disturbances from external sources, e.g. variation in patients’ body composition (especially weight) or additional clinical tasks (defibrillation, intra-venous or intra-osseous access, administration of medicine).

Conclusion

This simulation based study will provide important information on the quality and ease of use comparing three commercially available mechanical chest compression devices. This will provide important information for choosing the optimal system to use in the pre-hospital setting.

Participants

Department of Anaesthesia and Intensive Care, Aalborg University Hospital

Thomas Lass Klitgaard, MD PhD

Simon Ladehoff Thomsen, MD

Bodil Steen Rasmussen, MD PhD

Den Præhospitale Virksomhed, North Denmark Region

Jannie Kristine Bang Gram, Responsible for Development and Research, Sociologist PhD

Fritz Møller, Paramedic

Peter Brian Madsen Burholt, Paramedic

Jacob Thornvig Andersen, Paramedic

Niels Ole Saaby Hald, MD

Department of Clinical Medicine, Aalborg University

Mads Kjeldsen, BS

Rune Pedersen, BS

Conflicts of interest

The companies responsible for each chest compression device have no role in the design of this study nor in data collection, management, analysis, or interpretation. None of the above mentioned participants have any conflicts of interest in relation to this study.

Ethics

We will obtain written and oral informed consent for participating in this study from all study participants. They will at any time be able to withdraw from the study. All written personal identifiable information will be anonymised, and each participant will be assigned an individual identity code, e.g. PM1.

Dissemination

We will publish the results from this study in an international peer-reviewed medical journal, regardless of the findings. The decision to publish the collected data resides with the abovementioned participants.

References

1. 	Dansk Hjertestopregister Årsrapport 2022. 2022;(december):1–55.

2. 	Kleinman ME, Brennan EE, Goldberger ZD, et al. Part 5: Adult basic life support and cardiopulmonary resuscitation quality: 2015 American Heart Association guidelines update for cardiopulmonary resuscitation and emergency cardiovascular care. Circulation. 2015;132(18):S414–35. doi:10.1161/CIR.0000000000000259

3. 	Perkins GD, Graesner JT, Semeraro F, et al. European Resuscitation Council Guidelines 2021: Executive summary. Resuscitation. 2021;161:1–60. doi:10.1016/j.resuscitation.2021.02.003

4. 	Wang PL, Brooks SC. Mechanical versus manual chest compressions for cardiac arrest. Cochrane Database of Systematic Reviews. 2018;(8):CD007260. doi:10.1002/14651858.CD007260

5. 	Hallstrom A, Rea TD, Sayre MR, et al. Manual Chest Compression vs Use of an Automated Chest Compression Device During Resuscitation Following Out-of-Hospital Cardiac Arrest - A Randomized Trial. Jama. 2006;295(22):2620–8.

6. 	S M. Comparative Hemodynamic Evaluation of the LUCAS® Device and Manual Chest Compression in Patients with Out-of-Hospital Cardiac Arrest. Journal of Clinical Intensive Care and Medicine. 2017;2(1):016–24. doi:10.29328/journal.jcicm.1001005

7. 	Gässler H, Ventzke MM, Lampl L, et al. Transport with ongoing resuscitation: A comparison between manual and mechanical compression. Emergency Medicine Journal. 2013;30(7):589–92. doi:10.1136/emermed-2012-201142

8. 	Drinhaus H, Nüsgen S, Adams N, et al. Rescue under ongoing CPR from an upper floor: Evaluation of three different evacuation routes and mechanical and manual chest compressions: A manikin trial. Scandinavian Journal of Trauma, Resuscitation and Emergency Medicine. 2020;28(1):1–8. doi:10.1186/s13049-020-0709-0

9. 	Bekgöz B, Şan İ, Ergin M. Quality Comparison of the Manual Chest Compression and the Mechanical Chest Compression During Difficult Transport Conditions. Journal of Emergency Medicine. 2020;58(3):432–8. doi:10.1016/j.jemermed.2019.11.045

10. 	San I, Bekgöz B, Ergin M, et al. Manual cardiopulmonary resuscitation versus mechanical cardiopulmonary resuscitation: Which one is more effective during ambulance transport? Turkish Journal of Emergency Medicine. 2021;21(2):69–74. doi:10.4103/2452-2473.309135

11. 	Jörgens M, Königer J, Kanz KG, et al. Testing mechanical chest compression devices of different design for their suitability for prehospital patient transport - a simulator-based study. BMC Emergency Medicine. 2021;21(1):1–8. doi:10.1186/s12873-021-00409-3

12. 	Alexander E, Katharina T, Verena F, et al. Comparison of different mechanical chest compression devices in the alpine rescue setting: a randomized triple crossover experiment. Scandinavian journal of trauma, resuscitation and emergency medicine. 2021;29(1):84. doi:10.1186/s13049-021-00899-x

13. 	Kamauzaman THT, Hock JNT, Arithra A, et al. Simulation study on quality of cpr between manual chest compression and mechanical chest compression devices performed in ambulance. Medical Journal of Malaysia. 2021;76(2):171–6.

14. 	Venema AM, Sahinovic MM, Absalom AR, et al. Surviving the storm: manual vs. mechanical chest compressions onboard a lifeboat during bad weather conditions. Frontiers in Disaster and Emergency Medicine. 2023; 1. doi: 10.3389/femer.2023.1200857 doi:10.3389/femer.2023.1200857

15. 	Ong MEH, Anantharaman V. Out-of-hospital cardiac arrest: Manual or mechanical CPR? The Lancet. 2015;385(9972):920–2. doi:10.1016/S0140-6736(14)61941-3

16. 	Lee CA, Jung YS, Kim W. Comparison of Manual Versus Mechanical Chest Compression During Simulative out of Hospital Cardiac Arrest. Journal of The Korean Society of Emergency Medicine. 2012;23(4):486–92.

17. 	Lindblad P, Victorén AÅ, Axelsson C, et al. A Chest Compression Quality Evaluation Using Mechanical Chest Compressions under Different Working Situations in the Ambulance. International Journal of Clinical Medicine. 2015;06(08):530–7. doi:10.4236/ijcm.2015.68071

18. 	López J, Fernández SN, González R, et al. Comparison between manual and mechanical chest compressions during resuscitation in a pediatric animal model of asphyxial cardiac arrest. PLoS ONE. 2017;12(11):1–14. doi:10.1371/journal.pone.0188846

19. 	Eichhorn S, Mendoza Garcia A, Polski M, et al. Corpuls cpr resuscitation device generates superior emulated flows and pressures than LUCAS II in a mechanical thorax model. Australasian Physical and Engineering Sciences in Medicine. 2017;40(2):441–7. doi:10.1007/s13246-017-0537-3

20. 	Eichhorn S, Mendoza A, Prinzing A, et al. Corpuls CPR Generates Higher Mean Arterial Pressure Than LUCAS II in a Pig Model of Cardiac Arrest. BioMed Research International. 2017; 2017. doi: 10.1155/2017/5470406 doi:10.1155/2017/5470406

21. 	Milling L, Astrup BS, Mikkelsen S. Prehospital cardiopulmonary resuscitation with manual or mechanical chest compression: A study of compression-induced injuries. Acta Anaesthesiologica Scandinavica. 2019;63(6):789–95. doi:10.1111/aas.13347

22. 	Perkins GD, Lall R, Quinn T, et al. Mechanical versus manual chest compression for out-of-hospital cardiac arrest (PARAMEDIC): A pragmatic, cluster randomised controlled trial. The Lancet. 2015;385(9972):947–55. doi:10.1016/S0140-6736(14)61886-9

23. 	Khan SU, Lone AN, Talluri S, et al. Efficacy and safety of mechanical versus manual compression in cardiac arrest – A Bayesian network meta-analysis. Resuscitation. 2018;130(May):182–8. doi:10.1016/j.resuscitation.2018.05.005

24. 	Perkins GD, Graesner JT, Semeraro F, et al. European Resuscitation Council Guidelines 2021: Executive summary. Resuscitation. 2021;161:1–60. doi:10.1016/j.resuscitation.2021.02.003

25. 	Seewald S, Dopfer S, Wnent J, et al. Differences between manual CPR and corpuls cpr in regard to quality and outcome: study protocol of the comparing observational multi‐center prospective registry study on resuscitation (COMPRESS). Scandinavian Journal of Trauma, Resuscitation and Emergency Medicine. 2021;29(1):1–7. doi:10.1186/s13049-021-00855-9