## Slide 1: Racial bias in pulse oximetry: next steps
Dr Peter H. Charlton
University of Cambridge
City, University of London
Photo by Syed Ali on Unsplash
A presentation given on 22-May-2024 at an event hosted by the Institute of Physics and Engineering in Medicine (London, UK).
Available under a CC BY 4.0 Licence at https://doi.org/10.5281/zenodo.11267302

## Slide 2
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 3
Tackling racial bias in pulse oximetry
Photo by Dylan Gillis on Unsplash
Resources: Wealth of expertise across disciplines
Infrastructure: NHS ecosystem for research and implementation
Timeliness: Recent report and increasing use of pulse oximetry

## Slide 4
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 5
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 6
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 7: Guidance
Recommendation 1:
Equity in medical devices: independent review,,GOV.UK, 2024, https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
“Regulators, developers, manufacturers and healthcare professionals should take immediate mitigation actions to ensure existing pulse oximeter devices in the NHS can be used safely and equitably for all patient groups across the range of skin tones.”
Guidance for use
Looking at changes in readings
Using trends rather than thresholds to inform decisions
Educate clinicians on pulse oximeter technology
Manufacturers to update instructions for use
Developers and manufacturers to report performance across subgroups

## Slide 8
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 9: Regulation
Recommendation 2:
Equity in medical devices: independent review,,GOV.UK, 2024, https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
“The MHRA and approved bodies for medical devices should strengthen the standards for approval of new pulse oximeter devices to include sufficient clinical data to demonstrate accuracy overall and in groups with darker skin tones. Greater population representativeness in testing and calibration of devices should be stipulated.”
Ensure diversity in validity data:
  no. participants, range of skin tones, range of oxygenation levels
Validate in intended use population and setting
Comply with ISO standard
Make healthcare equity impact assessments an essential requirement

## Slide 10
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 11: Device development
Recommendation 3:
Equity in medical devices: independent review,,GOV.UK, 2024, https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
“Innovators, researchers and manufacturers should cooperate with public and patient participants to design better, smarter oximeters using innovative technologies to produce devices that are not biased by skin tone.”
Enhanced algorithms to address measurement bias
Multi-wavelength systems which measure and correct for skin pigmentation

## Slide 12
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 13
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 14: Guidance
Recommendation 1:
Equity in medical devices: independent review,,GOV.UK, 2024, https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
“Regulators, developers, manufacturers and healthcare professionals should take immediate mitigation actions to ensure existing pulse oximeter devices in the NHS can be used safely and equitably for all patient groups across the range of skin tones.”
Guidance for use
Looking at changes in readings
Using trends rather than thresholds to inform decisions
Educate clinicians on pulse oximeter technology
Manufacturers to update instructions for use
Developers and manufacturers to report performance across subgroups

## Slide 15: Using trends rather than thresholds to inform decisions
Royal College of Physicians, National Early Warning Score (NEWS) 2 Standardising the assessment of acute-illness severity in the NHS, 2017, https://www.rcp.ac.uk/improving-care/resources/national-early-warning-score-news-2/

## Slide 16: Using trends rather than thresholds to inform decisions
What if you don’t have a baseline for the trend? e.g.
  First measurement
  Patient is sick at first encounter
  What if pulse oximeters don’t monitor trends accurately?
Perkins GD et al., Critical Care, 2003, https://doi.org/10.1186/cc2339 (CC BY 4.0)

## Slide 17: Using trends rather than thresholds to inform decisions
What if you don’t have a baseline for the trend? e.g.
  First measurement
  Patient is sick at first encounter
  What if pulse oximeters don’t monitor trends accurately?
  What if the accuracy of trend monitoring differs with skin pigmentation?

## Slide 18
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 19: Device development
Recommendation 3:
Equity in medical devices: independent review,,GOV.UK, 2024, https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
“Innovators, researchers and manufacturers should cooperate with public and patient participants to design better, smarter oximeters using innovative technologies to produce devices that are not biased by skin tone.”
Enhanced algorithms to address measurement bias
Multi-wavelength systems which measure and correct for skin pigmentation

## Slide 20: Pulse oximeter algorithms
Reference blood oxygen saturation
Ratio of
ratios
Suggestion 1:
Pulse oximeters are inaccurate because their equations are derived from measurements on predominantly White individuals.
Increasing the diversity of training data would remove bias

## Slide 21: Pulse oximeter algorithms
Reference blood oxygen saturation
Ratio of
ratios
Suggestion 1:
Pulse oximeters are inaccurate because their equations are derived from measurements on predominantly White individuals.
Increasing the diversity of training data would remove bias
It might reduce bias, but wouldn’t eliminate it if the true relationship differs with skin tone.
Different skin tones

## Slide 22: Pulse oximeter algorithms
Reference blood oxygen saturation
Ratio of
ratios
Suggestion 2:
The relationships differ between skin tones.
We need to learn the different relationships, and have a way of identifying skin tone to select the appropriate relationship to use (e.g. multiwavelength PPG).
Perhaps.
Different skin tones

## Slide 23: Pulse oximeter algorithms
Reference blood oxygen saturation
Ratio of
ratios
Suggestion 3:
The relationship(s) are less strong for some skin tones when using traditional pulse oximetry sensing.
We need to identify improved sensing techniques (e.g. multiwavelength PPG).
Perhaps.
Different skin tones

## Slide 24
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement

## Slide 25: Potential research directions
Understanding current practice in the NHS
  What pulse oximeters are being used?
  How do they inform decision making?
Level of inaccuracy in pulse oximeters in UK
  Retrospective study using routine blood gas measurements as reference.
  Consider not only differences in accuracy between skin tones, but also overall accuracy and precision.
Estimates of precision vary between studies, methodologies, and devices.
e.g. Limits of agreement of +/- 4%
Blanchet M-A et al., Respiratory Care, 2023, https://doi.org/10.4187/respcare.10582

## Slide 26: Potential research directions
Understanding current practice in the NHS
  What pulse oximeters are being used?
  How do they inform decision making?
Level of inaccuracy in pulse oximeters in UK
  Retrospective study using routine blood gas measurements as reference.
  Consider not only differences in accuracy between skin tones, but also overall accuracy and precision.
  Consider differences between devices
“One oximeter overestimated SaO2 (Philips, +0.9%), whereas the 3 others underestimated SaO2 (Nonin −3.1%, Nellcor −0.3%, Masimo −0.2%).”
Blanchet M-A et al., Respiratory Care, 2023, https://doi.org/10.4187/respcare.10582

## Slide 27: Potential research directions
Understanding current practice in the NHS
  What pulse oximeters are being used?
  How do they inform decision making?
Level of inaccuracy in pulse oximeters in UK
  Retrospective study using routine blood gas measurements as reference.
  Consider not only differences in accuracy between skin tones, but also overall accuracy and precision.
  Consider differences between devices
  Consider additional populations (e.g. neonates)
Impact of inaccuracies in pulse oximetry on outcomes
  Assessing the scale of the problem in the UK
  Retrospective study (e.g. using ICNARC) on occult hypoxaemia, or morbidity and mortality
  Prospective study using objective quantification of skin tone
Mitigation of differential bias with existing devices
  Learn appropriate corrections
  (somehow) apply these in practice (e.g. based on self-reported ethnicity)

## Slide 28
Sonia Sevilla, https://commons.wikimedia.org/wiki/File:Kamar_Zard_Buzhan_-_Nishapur_1.jpg (CC0 1.0)
Much work has been done to recognise the problem of racial bias in pulse oximetry.
There is much work to be done to mitigate against and overcome this problem. Areas for interdisciplinary work include:
Providing guidance
Enhancing regulation
Device development
Research
Today we have a valuable opportunity to catalyse such work.
Notes: To conclude,
[click] There is great opportunity to use photoplethysmography in public health
[click] Clinical applications are emerging, including detecting atrial fibrillation
[click] And, when used clinically, photoplethysmgoraphy should be like a climbing rope. [click] It should be highly reliable, validated, and used for specified purposes.

## Slide 29
Sjoding MW et al., NEJM, 2020, https://doi.org/10.1056/NEJMc2029240
Photo by Syed Ali on Unsplash
Equity in medical devices: independent review,,GOV.UK, 2024,
https://www.gov.uk/government/publications/equity-in-medical-devices-independent-review-final-report (OGL 3.0)
Guidance
Device
development
Regulation
Research
New England Journal of Medicine
Racial bias in pulse oximetry measurement