FACS Aria II Shut Down

- Scope and Applicability: This protocol provides step by step instructions for shutting down the BD FACS Aria II and cleaning it of potentially hazardous materials.

- Materials:

  - MilliQ H2O

  - Sheath fluid. (BD Bioscience 342003)

  - Bleach

  - 70% Ethanol

  - Kimwipes

- Equipment:

  - P2, P10, P1000 pipets

  - 1 Liter graduated cylinder

  - BD Falcon 5 ml polystyrene round bottom tubes

  - FACSAria II machine

  - Biohazard disposal container

- Safety:

  - Nitrile Gloves

  - Eye protection

  - Lab coat

- Output:

  - Safely shutdown and cleaned BD FACS Aria II

- Reference Documents:

  - FACSDiva Software manual

http://static.bdbiosciences.com/documents/FACSDiva8-QRG-Administrator-Guide.pdf

  - FACS Aria II User Guide

http://static.bdbiosciences.com/documents/BD_FACSAria_II_User_Guide.pdf

Warning:  Personal Protective Equipment (PPE) should be used at all times while operating this protocol. If you are unsure what PPE you should be using, see your immediate supervisor.

- Methodology/Procedures:

  - Daily Shutdown Procedure.

    - Turn off the stream

    - Remove nozzle and insert the integrated closed loop nozzle

    - Clean Flow cell 2 times with MilliQ  water

      - Select Cytometer > Cleaning Modes > Clean Flow Cell.

      - Install a tube containing approximately 4 ml of MilliQH2O, then click Okay.

    - Turn off the Cytometer main power ( press big green button on the left side of the Aria).

    - Exit BD Coherent Connection software.

    - Exit Cytometer Tera Term Program.

    - Exit BD FACSDiva software and shut down the computer.

    - Depressurize the sheath tanks and refill with sheath fluid.

    - Empty waste in sink (add bleach if it was not already in the waste tank).

    - Turn off Aerosol Management System and Water Cooler if these were used.

  - Weekly Shutdown Procedure ( Fluidics Shutdown)

    - Turn off the stream.

    - Select Cytometer > Fluidics Shutdown. The following window opens:  MAKE SURE THERE IS PLENTY OF 70% ETOH IN THE TANK BEFORE STARTING (FILL TO THE WELD LINE)

    - Remove the nozzle from the flow cell assembly and click Done.

    - Insert the integrated closed loop nozzle into the flow cell assembly and click Done.

    - Disconnect the air line from the sheath tank and connect it to the air port on the ETON shutdown tank.

    - Disconnect the fluid line from the bottom side of the sheath filter and connect it to the ethanol filter on the ethanol shutdown tank.

    - Click Done.

    - When prompted, install a tube containing approximately 4 ml of MilliQH2Oon the loading port, then click Done.

    - Click Okay when you see a message that the system can be turned off.

    - Exit BD Coherent Connection software.

    - Exit Cytometer Tera Term Program..

    - Exit BD FACSDiva software and shut down the computer.

    - Depressurize the sheath tanks and refill with sheath fluid.

    - Empty waste in sink (add bleach if it was not already in the waste tank).

    - Turn off Aerosol Management System and Water Cooler if these were used.

  - Extra Cleaning

    - Load and run a tube of 10% bleach for at least a minute at a flow rate of 11.

    - Repeat with a tube of DI water.

    - Turn stream off.

    - Repeat step 2 through step12 from the Weekly Shutdown Procedure.

  - Monthly Aseptic Cleaning Procedure

    - In order to maintain sterility of the fluidic system on the FACS machines, we conduct an aseptic clean the first workday of every month.

    - For aseptic clean, run the “Prepare for Aseptic Sort” cleaning module.  This takes roughly 1 hour.

    - Press “Cytometer” on the file menu, then go to “Cleaning Modes”, and last “Prepare for Aseptic Sort”.

    - Follow wizard instruction to carry out aseptic clean.  Also consult BD FACSAria II User’s Guide pages 187-190.

- Appendix: Handling biologically hazardous samples

  - Aerosol Management System (AMS)

    - Our Aria has an Aerosol Management System that is used to contain all aerosols that arise during the sort process.  This consists of a separate vacuum unit that pushes air through a 0.12-µm ultra-low penetration air (ULPA) filter, and a series of drilled holes to allow airflow through the sort chamber and collection area.  All doors must be closed for this to work correctly.

    - If using this system, conduct all sorting setups with it on, since the airflow can affect the side streams.

    - When sorting human material, or other potentially hazardous samples, the Aerosol Management System (AMS) must be used.

    - Ensure that the main POWER button on the back of the evacuator is on.

    - Press the POWER button on the membrane panel of the evacuator.

    - Ensure that the suction control rate is set to 20%.(Do not set the suction control rate above 20%, higher rates could affect the stability of the side streams.)

    - Verify that the filter flow gauge reads less than 2.4 inches.

    - Sort as usual.

    - If a clog is encountered, do the following steps:

      - Turn stream off (keep top and flow chamber doors CLOSED)

      - Turn up suction rate to 100% and allow at least 1 minute to evacuate aerosols from the sort chamber.

      - Remove nozzle and conduct appropriate cleaning procedures. Place a nozzle in a clean FACS tube with MilliQH2O, close tube with a cap and sonicare it for 1 min in Sonicator. Take out nozzle from a tube and dry with kimwipes.

      - Reset suction to 20% and begin stream again.

      - Put AMS on standby by pressing the POWER button on the membrane panel when you are finished sorting.